## Slide 1: Progress Report Spirometry Software
Group 36
Abby Cohen
Team mates: Jessie Butts, Olga Neyman
Client: Dr. Klaesner

## Slide 2: Need
  80% of developing countries lack access to spirometry equipment – World Health Organization
  Factors like smoking and use of biomass fuels are increasing prevalence of respiratory diseases
Notes: Hindrances- cost, access to doctors, equipment needs to be calibrated, disposable mouthpieces

## Slide 3: The Big Picture
Spirometer Hardware
User’s Laptop
Server
Record Data
Troubleshooting
Data Processing
Data Analysis
Results

## Slide 4: Audio Sensor
 | Desired Criteria | Weight | Omnidirectional | Laptop | Unidirectional | Digital
Cost | < 5$ | 10 | 10 | 10 | 10 | 0
Frequency Response | (0-1kHz) | 10 | 9 | 9 | 9 | 9
Output Type | Analog | 10 | 10 | 10 | 10 | 0
Operating Temp | 0~40° C | 10 | 10 | 10 | 10 | 10
Signal-to-noise | > 60 dB | 8 | 10 | n/a | 10 | 9
Directionality | Unidirectional | 8 | 0 | 0 | 10 | 0
Shape | Circular | 8 | 10 | 10 | 10 | 0
Diameter | 0.16 cm– 1.27 cm | 8 | 10 | 0 | 10 | 0
Height | < 2.5 cm | 8 | 10 | 0 | 10 | 10
Impedance | < 600 Ω | 7 | 5 | n/a | 9 | n/a
Weight | < 10 g | 6 | 10 | n/a | 10 | 10
Current Supply | ≤ 500 μA | 6 | 10 | n/a | 10 | 0
Sensitivity | < -55 dB | 6 | 5 | 9 | 10 | 0
Total |  |  | 895 |  | 1033 | 
Notes: Limited initially… termination points, solder leads to different microphones
Correct specifications rather than just the termination point
Unidirectional-
Omnidirectional
Laptop
Digital
Omnidirection vs

## Slide 5: Microphone Omnidirectional vs. Unidirectional
Microphone diaphragm reaction depends of direction of sound
Omnidirectional
Unidirectional

## Slide 6: Audio Sensor
 | Desired Criteria | Weight | Omnidirectional | Laptop | Unidirectional | Digital
Cost | < 5$ | 10 | 10 | 10 | 10 | 0
Frequency Response | (0-1kHz) | 10 | 9 | 9 | 9 | 9
Output Type | Analog | 10 | 10 | 10 | 10 | 0
Operating Temp | 0~40° C | 10 | 10 | 10 | 10 | 10
Signal-to-noise | > 60 dB | 8 | 10 | n/a | 10 | 9
Directionality | Unidirectional | 8 | 0 | 0 | 10 | 0
Shape | Circular | 8 | 10 | 10 | 10 | 0
Diameter | 0.16 cm– 1.27 cm | 8 | 10 | 0 | 10 | 0
Height | < 2.5 cm | 8 | 10 | 0 | 10 | 10
Impedance | < 600 Ω | 7 | 5 | n/a | 9 | n/a
Weight | < 10 g | 6 | 10 | n/a | 10 | 10
Current Supply | ≤ 500 μA | 6 | 10 | n/a | 10 | 0
Sensitivity | < -55 dB | 6 | 5 | 9 | 10 | 0
Total |  |  | 895 |  | 1033 | 
Notes: Limited initially… termination points, solder leads to different microphones
Correct specifications rather than just the termination point
Unidirectional-
Omnidirectional
Laptop
Digital
Omnidirection vs

## Slide 7: Noise Cancellation Hardware Approach
 | Desired Criteria | Weight | Microphone Location | Microphone Number | Pop Filter Analog
Accuracy | +-3% | 10 | 7 | 8 | 6
Cost | <$5 | 10 | 10 | 4 | 8
Precision | +-5% | 8 | 8 | 6 | 10
Noise Reduction |  | 8 | 5 | 8 | 8
Response time for analysis | <10 seconds | 4 | 10 | 6 | 10
Totals |  |  | 314 | 292 | 324

## Slide 8: Noise Cancellation Microphone Location
Relationship between frequency
  F1 = 2 * F2
Analysis
  Reynold’s Number
  Boundary Layer
    Located in the obstacle valley
    Located in a side chamber

## Slide 9: Noise Cancellation Microphone Number
    Differential Noise Cancellation
    Average/Directed Average
Differential noise cancellation
Average
Directed average
  Low flow rates
Analysis
  Microphone adapter
  Accuracy

## Slide 10: Noise Cancellation Pop Filter
Plosives
  Max out microphone
  Buffer
Analysis
  Materials
  Degradation
    Typical Use of a Pop Filter

## Slide 11: Noise Cancellation Hardware Approach
 | Desired Criteria | Weight | Microphone Location | Microphone Number | Pop Filter Analog
Accuracy | +-3% | 10 | 7 | 8 | 6
Cost | <$5 | 10 | 10 | 4 | 8
Precision | +-5% | 8 | 8 | 6 | 10
Noise Reduction |  | 8 | 5 | 8 | 8
Response time for analysis | <10 seconds | 4 | 10 | 6 | 10
Totals |  |  | 314 | 292 | 324

## Slide 12: Noise Cancellation Software Approach
 | Desired Criteria | Weight | Window | Averaging Multiple FFT's | Moving Average Filter | Digital Low-Pass Filter | Suppression Function | Spectral Subtraction
Accuracy | ±3% | 10 | 6 | 8 | 6 | 6 | 5 | 7
Complexity | <120 seconds | 3 | 8 | 6 | 8 | 4 | 4 | 7
Noise Reduction |  | 10 | 5 | 6 | 7 | 8 | 6 | 5
Flexibility |  | 6 | 3 | 8 | 7 | 4 | 6 | 3
Totals |  |  | 151 | 197 | 197 | 171 | 153 | 148

## Slide 13: Noise Cancellation Windowing and Filters
Hanning Window
Moving Average Filter
  Reduces effects of peaks of random noise
Low Pass Filter
  Frequencies between 0-1000 Hz
Suppression Function
  Iterative approach to noise cancellation

## Slide 14: Noise Cancellation Averaging Multiple FFTs

## Slide 15: Noise Cancellation Spectral Subtraction

## Slide 16: Noise Cancellation Software Approach
 | Desired Criteria | Weight | Window | Averaging Multiple FFT's | Moving Average Filter | Digital Low-Pass Filter | Suppression Function | Spectral Subtraction
Accuracy | ±3% | 10 | 6 | 8 | 6 | 6 | 5 | 7
Complexity | <120 seconds | 3 | 8 | 6 | 8 | 4 | 4 | 7
Noise Reduction |  | 10 | 5 | 6 | 7 | 8 | 6 | 5
Flexibility |  | 6 | 3 | 8 | 7 | 4 | 6 | 3
Totals |  |  | 151 | 197 | 197 | 171 | 153 | 148

## Slide 17: Software Flow Chart
Spirometry Software
Training
Maneuver
Record Signal as WAV
Open Website
Patient Info
Download Recorder
Troubleshoot data
Noise Cancellation
Filtering and Windowing
Averaging
Suppression
STFT and Plot
Predicted Values
Data acceptability
Test Results

## Slide 18: Chosen Design to Implement
Hardware
  Electret Condenser Mic (CMI-5247TF-K)
  Pop filter for Noise cancellation
Software: Java platform with JSP framework
  Downloadable component: WAV file
  Web app component
    Troubleshooting
    Noise Cancellation
      Filters
      Windowing
      Iterative approach
      Averagin
Display
  Reliability of data
  Results showing predicted values, flow volume curves and derived values

## Slide 19: Updated Design Schedule
Tasks | 15-Oct | 22-Oct | 29-Oct | 5-Nov | 12-Nov | 19-Nov | 28-Nov | 5-Dec | 11-Dec
Design Safe |  |  |  |  |  |  |  |  | 
Progress Written Report |  |  |  |  |  |  |  |  | 
Progress Oral Report |  |  |  |  |  |  |  |  | 
Implementation of Predicted Values |  |  |  |  |  |  |  |  | 
Interface Creation and Implementation |  |  |  |  |  |  |  |  | 
Back-end Noise Cancellation |  |  |  |  |  |  |  |  | 
Development of Training Module |  |  |  |  |  |  |  |  | 
Web App Implementation |  |  |  |  |  |  |  |  | 
Troubleshooting Software |  |  |  |  |  |  |  |  | 
Final Oral Report |  |  |  |  |  |  |  |  | 
Final Written Report |  |  |  |  |  |  |  |  | 
Project Poster |  |  |  |  |  |  |  |  | 

## Slide 20: Updated Team Responsibilities
Team Member | Jessie | Olga | Abby
Completed Responsibilities | -Sensor analysis and selection | -First iteration of back-end analysis -Preliminary Oral Report | -First iteration of downloadable component -Design of Front End
Major Concerns Moving Forward | -Miscellaneous | -Web app creation | -Back end optimization
Detail Responsibilities | -Predicted Values -Design Safe -Training | -Interfacing between client and server | -Implementation of noise cancellation

## Slide 21: Questions?